GOVERNMENT OF INDIA • MINISTRY OF CONSUMER AFFAIRS, FOOD & PUBLIC DISTRIBUTION
DEPARTMENT OF CONSUMER AFFAIRS — DIRECTORATE OF LEGAL METROLOGY
NATIONAL LEGAL METROLOGY TYPE EVALUATION LABORATORY

PATTERN EVALUATION TEST REPORT
NON-AUTOMATIC WEIGHING INSTRUMENT (NAWI)

In accordance with OIML Recommendation R 76-1:2006 (E) & Legal Metrology Act, 2009

| Report Number: | OIML-R76-2026-IND-0008 | Application No: | APP-NAWI-2026-0008 |
| --- | --- | --- | --- |
| Testing Date: | 2026-09-29 | Overall Status: | FINALIZED |

1. Instrument Specifications & Technical Characteristics

| Manufacturer Name | Bharat Precision Metrology | Instrument Type | Non-Automatic Bench Scale |
| --- | --- | --- | --- |
| Manufacturer Address | Okhla Industrial Area, Phase III, New Delhi | Model Designation | BPM-30-ADV |
| Accuracy Class | Class III | Serial Number | SN-2026-00008 |
| Max Capacity (Max) | 30.0 kg | Scale Interval (d) | 0.01 kg |
| Min Capacity (Min) | 0.2 kg | Verification Interval (e) | 0.01 kg |
| Scale Intervals (n) | n = 3000 | Tare Subtractive Max | -30.0 kg |
| Operating Temperature | -10.0°C to +40.0°C | Power Supply | 230V AC (+10% / -15%), 50Hz |
| Load Cell / Sensor | None | Firmware / Checksum | v1.0.0 (None) |

2. Laboratory and Environmental Conditions (Clause A.3)

| Laboratory Name | National Legal Metrology Type Evaluation Laboratory | Ambient Temp | 23.5 °C |
| --- | --- | --- | --- |
| Accreditation | NABL ISO/IEC 17025 Accredited & OIML Issuing Authority | Relative Humidity | 52.0 % RH |
| Local Gravity (g) | 9.7803 m/s² | Atmospheric Pressure | 1013.2 hPa |
| Standard Weights | Class M1 & F2 Working Standards (Traceable to NPL India) | Testing Officer | Er. Rajesh Sharma (Senior Metrological Officer) |

3. Weighing Performance Test Observations (Clause A.4.4)

| # | Dir | Load (kg) | Indication I | ΔL | Error Ec | Ec (e) | MPE (±e) | MPE (kg) | Status |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | INCR | 0 | 0 | 0.005 | +0.0000 | +0.00 e | ±0.5 e | ±0.0025 | PASS |
| 2 | INCR | 10 | 10 | 0.005 | +0.0000 | +0.00 e | ±1.0 e | ±0.0050 | PASS |
| 3 | INCR | 30 | 30 | 0.005 | +0.0000 | +0.00 e | ±1.5 e | ±0.0075 | PASS |
| 4 | DECR | 10 | 10 | 0.005 | +0.0000 | +0.00 e | ±1.0 e | ±0.0050 | PASS |
| 5 | DECR | 0 | 0 | 0.005 | +0.0000 | +0.00 e | ±0.5 e | ±0.0025 | PASS |

4. Prescribed Tests Compliance Summary

| Clause | Test Procedure | Prescribed Standard Limit | Verdict |
| --- | --- | --- | --- |
| Clause A.4.4 | Weighing Performance Test | Error ¦Ec¦ <= MPE across range; Hysteresis <= ¦MPE¦ | PASS |
| Clause A.4.10 | Repeatability Test (at 0.5 Max & Max) | Max difference between weighings <= ¦MPE¦ | PASS |
| Clause A.4.7 | Eccentricity / Off-Center Loading | ¦Ec¦ <= MPE at 1/3 Max at 4 corners and center | PASS |
| Clause A.4.2 | Zero-Setting and Zero-Tracking | Residual zero error <= +/- 0.25 e | PASS |
| Clause A.4.8 | Discrimination Test | Extra load 1.4 d gives noticeable change of indication | PASS |
| Clause A.5.3 | Static Temperatures (-10°C to +40°C) | Span error <= MPE; Zero drift <= 1 e / 5°C | PASS |
| Clause A.5.4 | Voltage Variations (+10% / -15%) | Indications remain within MPE limits | PASS |

STATUTORY MODEL APPROVAL VERDICT: COMPLIANT & APPROVED FOR USE

| TESTED & EVALUATED BY: Er. Rajesh Sharma (Senior Metrological Officer) Senior Metrological Officer, Legal Metrology | VERIFIED & APPROVED BY: Dr. Priya V. Iyer (Director of Legal Metrology) Director of Legal Metrology, Government of India |
| --- | --- |
| Digital Sign Hash: SHA256-AUTHENTICATED... Date: 2026-09-29 10:31:46 | [DIGITALLY SIGNED & VERIFIED] Official Statutory Repository, Legal Metrology Act, 2009 |