GOVERNMENT OF INDIA • MINISTRY OF CONSUMER AFFAIRS, FOOD & PUBLIC DISTRIBUTION
DEPARTMENT OF CONSUMER AFFAIRS — DIRECTORATE OF LEGAL METROLOGY
NATIONAL LEGAL METROLOGY TYPE EVALUATION LABORATORY

PATTERN EVALUATION TEST REPORT
NON-AUTOMATIC WEIGHING INSTRUMENT (NAWI)

In accordance with OIML Recommendation R 76-1:2006 (E) & Legal Metrology Act, 2009

| Report Number: | OIML-R76-2026-IND-0027 | Application No: | APP-NAWI-2026-0027 |
| --- | --- | --- | --- |
| Testing Date: | 2026-09-20 | Overall Status: | TESTING IN PROGRESS |

1. Instrument Specifications & Technical Characteristics

| Manufacturer Name | Bharat Heavy Weighbridges & Sensorics Ltd | Instrument Type | Pitless Electronic Weighbridge |
| --- | --- | --- | --- |
| Manufacturer Address | Industrial Area Phase 1, Chandigarh - 160002 | Model Designation | BHW-Titan-60T |
| Accuracy Class | Class III | Serial Number | SN-2026-WB-60T-004 |
| Max Capacity (Max) | 60000.0 kg | Scale Interval (d) | 20.0 kg |
| Min Capacity (Min) | 400.0 kg | Verification Interval (e) | 20.0 kg |
| Scale Intervals (n) | n = 3000 | Tare Subtractive Max | -60000.0 kg |
| Operating Temperature | -10.0°C to +50.0°C | Power Supply | 230V AC 50Hz with 12V Heavy Industrial UPS |
| Load Cell / Sensor | 6 x Flintec RC3 30t Stainless Steel Canister Load Cells (OIML R60 C3) | Firmware / Checksum | v4.0.2-WB (CRC: 0xF728D1B0) |

2. Laboratory and Environmental Conditions (Clause A.3)

| Laboratory Name | Regional Legal Metrology Verification Center (Field Station) | Ambient Temp | 29.0 °C |
| --- | --- | --- | --- |
| Accreditation | NABL ISO/IEC 17025 Accredited | Relative Humidity | 58.0 % RH |
| Local Gravity (g) | 9.7803 m/s² | Atmospheric Pressure | 1008.5 hPa |
| Standard Weights | Mobile Test Van with 20 tonnes of Calibrated Class M1 Block Weights | Testing Officer | Er. Rajesh Sharma (Senior Metrological Officer) |

3. Weighing Performance Test Observations (Clause A.4.4)

| # | Dir | Load (kg) | Indication I | ΔL | Error Ec | Ec (e) | MPE (±e) | MPE (kg) | Status |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | INCR | 0 | 0 | 5 | +0.0000 | +0.00 e | ±0.5 e | ±0.0025 | PASS |
| 2 | INCR | 1000 | 1000 | 5 | +0.0000 | +0.00 e | ±1.5 e | ±0.0075 | PASS |
| 3 | INCR | 10000 | 10000 | 6 | -1.0000 | -200.00 e | ±1.5 e | ±0.0075 | FAIL |
| 4 | INCR | 20000 | 20000 | 5 | +0.0000 | +0.00 e | ±1.5 e | ±0.0075 | PASS |
| 5 | INCR | 40000 | 40000 | 8 | -3.0000 | -600.00 e | ±1.5 e | ±0.0075 | FAIL |
| 6 | INCR | 60000 | 60000 | 5 | +0.0000 | +0.00 e | ±1.5 e | ±0.0075 | PASS |

4. Prescribed Tests Compliance Summary

| Clause | Test Procedure | Prescribed Standard Limit | Verdict |
| --- | --- | --- | --- |
| Clause A.4.4 | Weighing Performance Test | Error ¦Ec¦ <= MPE across range; Hysteresis <= ¦MPE¦ | FAIL |
| Clause A.4.10 | Repeatability Test (at 0.5 Max & Max) | Max difference between weighings <= ¦MPE¦ | PENDING |
| Clause A.4.7 | Eccentricity / Off-Center Loading | ¦Ec¦ <= MPE at 1/3 Max at 4 corners and center | FAIL |
| Clause A.4.2 | Zero-Setting and Zero-Tracking | Residual zero error <= +/- 0.25 e | PENDING |
| Clause A.4.8 | Discrimination Test | Extra load 1.4 d gives noticeable change of indication | PENDING |
| Clause A.5.3 | Static Temperatures (-10°C to +40°C) | Span error <= MPE; Zero drift <= 1 e / 5°C | PENDING |
| Clause A.5.4 | Voltage Variations (+10% / -15%) | Indications remain within MPE limits | PENDING |

STATUTORY MODEL APPROVAL VERDICT: NON-COMPLIANT & REJECTED

| TESTED & EVALUATED BY: Er. Rajesh Sharma (Senior Metrological Officer) Senior Metrological Officer, Legal Metrology | VERIFIED & APPROVED BY: Dr. Priya V. Iyer (Director of Legal Metrology) Director of Legal Metrology, Government of India |
| --- | --- |
| Digital Sign Hash: 4b556a47d271722d89978be196a1f499... Date: 2026-09-28T10:25:07.609452+00:00 | [DIGITALLY SIGNED & VERIFIED] Official Statutory Repository, Legal Metrology Act, 2009 |